ISS Report 1/29/15

All references are according to D1101059-v5

All filters have been removed for shipping.

AEI

S/N S1202970

* Missing peek connectors at the top to protect quad pins.
* Missing part #37 (beveled counter-bored washer under QPD)
* No strain relief installed.
* LHO height PD Base D1101074

S/N S1202972

* Missing part #37.
* No strain relief.
* LLO height PD Base D1101073.

LLO

S/N S1202965

* Missing peek connectors at the top to protect quad pins.
* Missing part #27x(2) and part #28x(2)
* Needs an alignment (QPD was loose)

S/N S1202967

* Missing part #37

LHO

S/N S1202968

* ISS Array Cover was switched to D1300720 s/n 008 (had an older version that was flappy)
* Missing part #37

S/N S1202969

* Missing peek connectors at the top to protect quad pins.
* Missing part #37

Note: #37 that I refer to (D1400146 beveled counter-bored washer under QPD) is available (6 in all) but the alignment of the quad was done without it installed. Is it necessary? It is installed on S1202965 only because the QPD was loose so I decided to do a fit check.