## Sheet1
ETMY
Upper Segment
Measured Current(milliamps) | Input Current(milliamps) | Voltage (volts) | Notes | Using Pin 3 and 16 on the 25 pin connection on the back of the Segment Driver, 47 ohms measured at the beginning of test
-0.31 | 0 | 0.008
99.2 | 100 | 1.732
198.1 | 200 | 3.455
297 | 300 | 5.18
395.9 | 400 | 6.91
494.8 | 500 | 8.63
593.6 | 600 | 10.37
690.5 | 700 | 12.07
785.4 | 800 | 13.75
811.1 | 900 | 14.22
811.2 | 1000 | 14.23
Lower Segment
Measured Current(milliamps) | Input Current(milliamps) | Voltage (volts) | Notes | Using Pin 9 and 22 on the 25 pin connection on the back of the Segment Driver, 48 ohms measured at the beginning of test.
1.6 | 0 | 0.035
171.92 | 100 | 2.967
271 | 200 | 4.683
370.2 | 300 | 6.4
469.3 | 400 | 8.12
568.5 | 500 | 9.84
667.6 | 600 | 11.57
764.7 | 700 | 13.25
860 | 800 | 14.95
885.8 | 900 | 15.41
885.8 | 1000 | 15.42
ITMY
Upper Segment
Measured Current(milliamps) | Input Current(milliamps) | Voltage (volts) | Notes | Using Pin 3 and 16 on the 25 pin connection on the back of the Segment Driver, 44.6 ohms measured at the beginning of test. Gains are set differently than ETMY
0.018 | 0 | 0.002
98.2 | 100 | 3.522
196.48 | 200 | 7.05
295.7 | 300 | 10.58
392.93 | 400 | 13.13
491.18 | 500 | 17.7
589.48 | 600 | 21.31
685.02 | 700 | 24.93
737 | 800 | 27.19
732.9 | 900 | 27.23
733 | 1000 | 27.24
Lower Segment
Measured Current(milliamps) | Input Current(milliamps) | Voltage (volts) | Notes | Using Pin 9 and 22 on the 25 pin connection on the back of the Segment Driver, measured 44.6 ohms when the driver was turned off but it jumped up to 70.1 ohms when the driver was turned on
0.021 | 0 | 0.026
97.25 | 100 | 3.476
195.05 | 200 | 6.98
292.85 | 300 | 10.49
390.67 | 400 | 14.07
488.48 | 500 | 17.58
586.41 | 600 | 21.14
681.6 | 700 | 24.66
745.5 | 800 | 27.16
742.9 | 900 | 27.18
740 | 1000 | 27.18
Upper Segment
Measured Current(milliamps) | Input Current(milliamps) | Voltage (volts) | Notes | Using Pin 3 and 16 on the 25 pin connection on the back of the Segment Driver. Using matching gain values as ETMY.
0.002 | 0 | 0.03
193.27 | 100 | 3.44
386.4 | 200 | 6.89
579.66 | 300 | 10.34
772.88 | 400 | 13.81
966.15 | 500 | 17.29
Lower Segment
Measured Current(milliamps) | Input Current(milliamps) | Voltage (volts) | Notes | Using Pin 9 and 22 on the 25 pin connection on the back of the Segment Driver.Using matching gain values as ETMY.
0.045 | 0 | 0.026
192.43 | 100 | 3.411
385.03 | 200 | 6.85
579.58 | 300 | 10.3
773.13 | 400 | 12.75
966.72 | 500 | 17.21