## LHO End X
Laser SN01 | Incoming laser power from T1300100
Date | 2013-06-04 | Aug. 3, 2018 | Mar. 19, 2019 | Jan-18-2024 | April-22-2025
Laser Shutter Check |  | PASS | PASS | Pass | Pass
Max OFS Offset |  | 9.9 V | 7.8 V | 7.7V | 7.7V
95% OFS Offset |  | 9.405 V | 7.4 V | 7.3V | 7.315
Operating OFS Offset |  | 4.7025 V | 3.7 V | 3.85V | 3.6575
Laser Output Power | 2.05 W | 1.72 W | 1.75 W | 1.67W | 1.7
After-Laser Rejected Power |  | 56.5 mW | 54 mW | 36.6mW | 34.6mW
AOM Input Power |  | 1.60 W | 1.69 W | 1.59W | 1.65
Max Diffracted Power |  | 0.967 W | 1.13 W | 1.1W | 1.14
Un-Diffracted Power |  | 0.359 W | 0.306 W | 266mW | 0.307W
AOM Diffraction Efficiency |  | 60.44% | 67% | 69% | 0.69091
After-AOM Rejected Power |  | 12.8 mW | 13.7 mW | 14.7mW | 15.2mW
TxPD Power |  | 11.1 mW | 6.06 mW | 6.36mW | 11.9mW
OFSPD Power |  | 5.7 mW | 6.56 mW | 6.36mW | 5.7mW
Outer Beam Power |  | 228 mW | 0.257 W | 265mW | 0.536
Inner Beam Power |  | 228 mW | 0.255 W | 260mW | 0.526
Output Beam Power Ratio |  | 1 | 0.99222 | 0.9811 | 0.98134
OFS Gain |  | 38.19 dB | 41 dB | 39.6db | 39.6
OFS Phase Margin |  | 60.9 degrees | 55 degrees | 43deg | 42.1
ALOG |  |  |  | 75484

## LHO End Y
Laser SN08 | Incoming laser power from T1300100
Date | 2013-06-04 | Aug. 3, 2018 | Mar. 19, 2019 | Jan, 23, 2024 | May , 21,2025
Laser Shutter Check |  | PASS | PASS | Pass | pass
Max OFS Offset |  | 7.9 V | 5.75 V | 8.1V | 8
95% OFS Offset |  | 7.505 V | 5.45 V | 7.7 | 7.6
Operating OFS Offset |  | 3.7525 V | 2.725 V | 3.85 | 3.8
Laser Output Power | 2.03 W | 1.94 W | 1.92 W | 1.93W | 1.94 W
After-Laser Rejected Power |  | 5.13 mW | 0.383 mW | 4.1mW | 4.11 mW
AOM Input Power |  | 1.84 W | 1.86 W | 1.91W | 1.91 W
Max Diffracted Power |  | 1.55 W | 1.05 W | 1.60W | 1.60 W
Un-Diffracted Power |  | 256 mW | 0.74 W | 160mW | 0.285 W
AOM Diffraction Efficiency |  | 84.24% | 49% | 83.8% | 83.77%
After-AOM Rejected Power |  | 16.4 mW | 15.8 mW | 16.5mW | 15.5 mW
TxPD Power |  | 6.39 mW | 4.05 mW | 14.2mW | 6.51mW
OFSPD Power |  | 6.5 mW | 4.24mW | 6.68mw | 6.64 mW
Outer Beam Power |  | 0.716 W | 0.25 W | 0.773W | 0.355 W
Inner Beam Power |  | 0.714 W | 0.248 W | 0.782W | 0.352 W
Output Beam Power Ratio |  | 0.9972 | 0.992 | 0.9885 | 0.99155
OFS Gain |  | 38.20 dB | 43 dB | 50.6 | 38.5
OFS Phase Margin |  | 55 degrees | 56 degrees | 61.7deg | 58

## LLO End X
Laser SN06 | Incoming laser power from T1300100
Date | 2013-06-04
Laser Shutter Check
Max OFS Offset
95% OFS Offset
Operating OFS Offset
Laser Output Power | 2.13 W
Rejected Laser Power
AOM Input Power
Max Diffracted Power
Un-Diffracted Power
AOM Diffraction Efficiency
AOM Rejected Power
TxPD Power
OFSPD Power
Outer Beam Power
Inner Beam Power
Output Beam Power Ratio
OFS Gain
OFS Phase Margin

## LLO End Y
Laser SN07 | Incoming laser power from T1300100
Date | 2013-06-04
Laser Shutter Check
Max OFS Offset
95% OFS Offset
Operating OFS Offset
Laser Output Power | 2.11 W
Rejected Laser Power
AOM Input Power
Max Diffracted Power
Un-Diffracted Power
AOM Diffraction Efficiency
AOM Rejected Power
TxPD Power
OFSPD Power
Outer Beam Power
Inner Beam Power
Output Beam Power Ratio
OFS Gain
OFS Phase Margin

## LHO Pcal lab test setup
Laser SN05 | Incoming laser power from T1300100
Date | 2013-06-04 | Jan-17-2024
Laser Shutter Check |  | Pass
Max OFS Offset |  | 7.8V
95% OFS Offset |  | 7.4V
Operating OFS Offset |  | 3.7V
Laser Output Power |  | 1.98W
Rejected Laser Power |  | 15mW
AOM Input Power |  | 1.92W
Max Diffracted Power |  | 1.32W
Un-Diffracted Power |  | 440mW (-1 order 90mW)
AOM Diffraction Efficiency |  | 69%
AOM Rejected Power |  | 21.3mW
TxPD Power |  | 7.67mW
OFSPD Power |  | 9.43mW
Outer Beam Power |  | 318mw
Inner Beam Power |  | 310mW
Output Beam Power Ratio |  | 0.975
OFS Gain |  | 7.39
OFS Phase Margin |  | -136.4 Deg
ALOG |  | 75452